## Annex No. 2

## Technical Specification

### Purchase of Industrial Laboratory Hot Water Bath

| Parameter | Requirement |
| --- | --- |
| THE PARTICIPANT shall provide information or proof against each requirement. |  |
| Quantity | 1 pc. – Laboratory Hot Water Bath (High-Volume). |
| Purpose | Designed for reliable and accurate temperature control applications in laboratory, research, and quality control settings, such as sample warming, thawing, or temperature testing of materials/products. |
| 1. Capacity and Dimensions |  |
| * Total Filling Volume | Must accommodate volumes between 3 L – 26 L. |
| * Usable Immersion Depth | Minimum immersion depth: 17 cm. |
| * External Dimensions (W x D x H) | 630 mm – 650 mm (W) 340 mm – 360 mm (D) 250 mm – 270 mm (H). |
| * Nominal Weight | 12 kg. |
| 2. Performance and Temperature Control |  |
| * Working Temperature Range | +15C to +100C. |
| * Temperature Stability | Must maintain stability of 0.2C |
| * Heating Capacity | 1.8kW – 2.2kW. |
| * Ambient Operation | Suitable for ambient temperatures up to +40C. |
| 3. Features and Hygiene |  |
| * Construction Material | High-quality material (e.g., stainless steel, polymer) designed for durability and low maintenance. |
| * Drainage | Must include an integrated drain screw for quick, clean, and effortless emptying. |
| * Safety | Integrated protection against running dry. Mains switch protected against spray water. |
| * Interface | Bright display (e.g., LED) for easy reading. Must include an integrated timer function. |
| 4. Electrical Requirements |  |
| * Power Supply | 230 V, 50 Hz (Standard European single-phase). |
| * Power Consumption | Max. power consumption 9 A |
| Complete Functional Set (Mandatory) | The Supplier must ensure the delivery of a single fully functional unit, including the bath vessel, heating element, integrated timer, safety features, and a platform included as standard. The offer must explicitly include: installation materials, commissioning, and staff training. |
| Documentation | To be supplied in Latvian or English: user manual, electric diagram, maintenance instructions, and CE Declaration of Conformity. |
| Guarantee | Minimum 24 months full manufacturer or supplier warranty. |