
Annex No. 2
Technical Specification
Purchase of Industrial Laboratory Hot Water Bath

	Parameter
	Requirement

	THE PARTICIPANT shall provide information or proof against each requirement.
	

	Quantity
	1 pc. – Laboratory Hot Water Bath (High-Volume).

	Purpose
	Designed for reliable and accurate temperature control applications in laboratory, research, and quality control settings, such as sample warming, thawing, or temperature testing of materials/products.

	1. Capacity and Dimensions
	

	* Total Filling Volume
	Must accommodate volumes between 3 L – 26 L.

	* Usable Immersion Depth
	Minimum immersion depth: 17 cm.

	* External Dimensions (W x D x H)
	630 mm – 650 mm (W) 
340 mm – 360 mm (D) 
250 mm – 270 mm (H).

	* Nominal Weight
	12 kg.

	2. Performance and Temperature Control
	

	* Working Temperature Range
	+15C to +100C.

	* Temperature Stability
	Must maintain stability of 0.2C 

	* Heating Capacity
	1.8kW – 2.2kW.

	* Ambient Operation
	Suitable for ambient temperatures up to +40C.

	3. Features and Hygiene
	

	* Construction Material
	High-quality material (e.g., stainless steel, polymer) designed for durability and low maintenance.

	* Drainage
	Must include an integrated drain screw for quick, clean, and effortless emptying.

	* Safety
	Integrated protection against running dry. Mains switch protected against spray water.

	* Interface
	Bright display (e.g., LED) for easy reading. Must include an integrated timer function.

	4. Electrical Requirements
	

	* Power Supply
	230 V, 50 Hz (Standard European single-phase).

	* Power Consumption
	Max. power consumption 9 A

	Complete Functional Set (Mandatory)
	The Supplier must ensure the delivery of a single fully functional unit, including the bath vessel, heating element, integrated timer, safety features, and a platform included as standard. The offer must explicitly include: installation materials, commissioning, and staff training.

	Documentation
	To be supplied in Latvian or English: user manual, electric diagram, maintenance instructions, and CE Declaration of Conformity.

	Guarantee
	Minimum 24 months full manufacturer or supplier warranty.



