## Annex No. 2

## Technical Specification

### Purchase of Industrial Refrigerated / Heating Circulator (High Volume)

| Parameter | Requirement |
| --- | --- |
| THE PARTICIPANT shall provide information or proof against each requirement. |  |
| Quantity | 1 pc. – Industrial Refrigerated / Heating Circulator. |
| Purpose | Designed for highly reliable and precise temperature control of external systems or larger samples/vessels placed in the unit's bath. Suitable for demanding research, quality control, and industrial tempering applications. |
| 1. Performance and Temperature Control |  |
| * Working Temperature Range | Must cover a broad range: -35C to +155C. |
| * Temperature Stability | Must achieve high temperature stability of 0.05C. |
| * Heating Capacity | Minimum heating capacity: 1.8 kW – 2.2 kW. |
| * Cooling Capacity | Minimum cooling capacity: 0.40 kW |
| * Refrigerant | Must use an environmentally friendly refrigerant (e.g., R290 or equivalent low GWP type). |
| 2. Bath and Pumping System |  |
| * Filling Volume | Internal bath vessel volume: 18 Litres – 28 Litres. |
| * Pumping Function | Suitable for internal bath applications and external applications. |
| * Pump Flow Rate | Max. pump flow rate: 14 L min – 16 L/min. |
| * Pump Pressure | Max. pump pressure: 0.3 bar – 0.4 bar. |
| * Drain | Must include a drain tap for easy fluid changes. |
| 3. Dimensions and Features |  |
| * External Dimensions (W x L x H) | Width 35 cm - 40 cm Length 55 cm - 60 cm Height 65 cm - 70 cm. |
| * Nominal Weight | Weight 38 kg - 41 kg. |
| * Connectivity | Must include USB connection for data logging and external remote control. |
| * Display | Bright, easy-to-read display. |
| 4. Electrical Requirements |  |
| * Supply Voltage | 230 V, 50 Hz |
| Complete Functional Set (Mandatory) | The Supplier must ensure the delivery of a single fully functional circulator unit, including the bath vessel, drain tap, integrated pump connections, and necessary operating accessories. |
| Documentation | To be supplied in Latvian or English: user manual, electric diagram, maintenance instructions, and CE Declaration of Conformity. |
| Guarantee | Minimum 24 months full manufacturer or supplier warranty. |